CVD SYSTEM: LEAK TEST REPORT

(July 15 & July 16 2018)

(20180715_184822.csv & 20180716_114209.csv) (Leak test 1 & Leak Test 2.py)

TEST 1:  A leak test was conducted with the following steps:

1. Evacuation & flush for 1 min each-- Twice

2. Pump down for 5 min (0.2 mbar) and then the butterfly valve was closed (basically a vacuum sealed chamber).

3. All zones were heated to 700 C for 30 minutes.

4. Hold at 700 C for 10 minutes.

5. End of Recipe & system was left for natural cooling.

Observations: A reddish brown colored film can be seen on cupper sheet. Dr. Sandeep will be doing XRD for further confirmation.

TEST 2:

# A second test has been conducted in a nitrogenous environment to compare both the results. The steps are:

1. Evacuation and flush

2. Evacuation for 1 minute

3. Flush for 5 minutes to make a pressure more that atm then gas flow was reduced to 25 sccm and purge valve was opened.

4. All zones were heated to 700 C for 30 minutes and  held for 10 minutes with standard gas flow of 25 sccm.

5. System was left for cooling with a gas flow of 25 sccm.

The system will cool down by evening and characterization will be done tomorrow ( XRD).